## Slide 1
Lectures:
Prof. Janne Ruokolainen, Dr. Hua Jiang, Dr. Jani Seitsonen, Dr. Ramzy Abdelaziz, Prof. Peter Liljeroth, Dr. Lide Yao
Schedule
17. 1. Introduction & Nanomicroscopy center (JR)
24. 1. TEM Basics (JR)
31. 1. Advanced TEM 1 (Hua)
7. 2. Advanced TEM 2 (Hua)
14. 2. Advanced TEM 3 (Hua)
21. 2. no lecture (exam period at Aalto)
28. 2. Cryo-TEM and Soft matter Sample preparation (JR)
7. 3. 3D-TEM-Tomography (Jani)
14. 3. STM/AFM (Peter)
21. 3. SEM (Ramzy)
28. 3. FIB and Sample preparation (Lide)
Summary:
Intro, Basic TEM, Cryo TEM ~ 3 lectures
Advanced TEM 2-3 lectures
(High resolution TEM and STEM, diffraction, spectroscopy EDX, EELS)
STM/AFM
FIB/sample preparation
SEM
Tomography
Additional Literature: (optional)
Book 1: Transmission electron microscopy Basics I (David William and Barry Carter) 2nd edition
Book 2: G.H. Michler ”Electron microscopy of polymers” (TEM, SEM, AFM..)
Book 3: A practical Guide to Tramsmission Electron Microscopy (Zhiping Luo)

## Slide 2: Cryo-Transmission electron microscopy: Sample preparation and material science applications
Prof. Janne Ruokolainen Department of Applied Physics

## Slide 3: Why CRYO?
Electron microscope column is in vacuum
Traditionally aqueous (water-containing) EM samples have been dehydrated
Cryo-EM allows electron microscopy of aqueous specimen in their natural, hydrated state (also other solvents are possible for materials science applications: gels, polymeric vesicles and micels)
Reducing beam damage for all materials..

## Slide 4: Beam damage
I) Radiolysis: Inelastic scattering (mainly electron-electron interactions such as ionization) breaks the chemical bonds of certain materials such as polymers and alkali halides. Polymers: Chain scission (produce low molecular weight compounds) or Cross-linking.
II) Knock-on damage (direct displacement of atoms..) Knock-on damage or sputtering: Knock-on damage is the displacement of atoms from the crystal lattice and creates point defects. If atoms are ejected from the specimen surface we call it sputtering. These processes are ubiquitous if the beam energy (E0) is high enough.
III) Heating: Phonons heat your specimen and heat is a major source of damage to polymers and biological tissue.
If thermal conduction is very high, heating is negligible.
If thermal conduction is poor, heating can be quite substantial
So, beam heating for metals is usually minimal but
small ceramic particles may be heated by the beam to
temperatures of 1700oC.

## Slide 5
How to minimize…
1. Low dose TEM: - As small intensity as possible
Small spot size, small cond. aperture, spread the beam, low magnification
Focusing etc. in the adjacent area or the specimen - Automatic low dose software
2. High voltage TEM
+ Electron mean free path longer when V increased  less inelastic scattering  less damage
knock-on damage is increased above 200 - 300kV
Less contrast in high voltages
However carbon nanotubes quite low voltage 80kV is a better ..
3. Cryo microscopy
At low temperatures degradation slower (either liquid nitrogen or liquid Helium cooling) (Dose typically ~10 electrons/Å^2 or some polymers even much less –but then the signal to noise ratio is a problem…)
4. Sample coating (thin film of carbon film), Replica …

## Slide 6: Nanomicroscopy center cryo-TEM’s
One Dedicated cryo-TEM:
Liquid nitrogen or liquid helium
+ Other TEM’s with cryo-Holders:
Single tilt cryo Holder
Multible specimen cryo holder
Cryo-rotation holder

## Slide 7: Nobel Prize for Chemistry 2017 for "developing cryo-electron microscopy for the high-resolution structure determination of biomolecules in solution".
The resolution progression of cryo-EM, illustrated by a representation of glutamate dehydroge-nase with an increasing level of detail from left to right. For a protein of this size, 334 kDa, the 1.8 Å resolution to the right (38) could only be achieved after 2012/13.
38. Merk, A., et al. (2016) Breaking cryo-EM resolution barriers to facilitate drug discovery. Cell 165, 1698-1707
DED = Direct Electron Detector Camera

## Slide 8: Sample preparation for cryo-TEM?

## Slide 9
(From Andres Kaech University of Zurich)

## Slide 11
- Problems and Artefacts

## Slide 12: How about just drying the samples..

## Slide 15: Critical point drying

## Slide 16: Critical point drying

## Slide 17
For TEM critical point drying is good for ”thin” samples like viruses etc.

## Slide 18
(From Andres Kaech University of Zurich)

## Slide 19: Cryo sample preparation methods
1) Bulk samples? (> 500 nm) (i.e too thick for TEM )
Cells, tissue, bulk materials, viscous bulk hydrogels,
2) Already ”thin” samples: e.g. small particles? (<500 nm, preferably < 200 nm)
Viruses, nanofibers in solution (CNF, CNC), liposomes, vesicles, micels, nanoparticles in solution, hydrogels
Freeze and then prepare thin sample for TEM
(example: cryo-microtoming thin sections )
Freeze and then just transfer directly to the TEM
Handbook of Cryo-Preparation Methods for Electron Microscopy, 2008
Edited by Annie Cavalier, Daniele Spehner, Bruno M. Humbel

## Slide 20
1. Bulk samples: High pressure freezing
High pressure (2000 bars) delays ice crystal nucleation
Sample thickness can be up to 200 micrometer

## Slide 21: High pressure freezing
Sample thickness can be up to 200 micrometers

## Slide 22: High pressure freezing: Protocol (Handbook of Cryo-Preparation Methods ..)

## Slide 23: Cryo sectioning and section pick-up and direct mounting on grid

## Slide 24: 2. Thin samples, small particles? Viruses, liposomes, vesicles, micels, solutions, hydrogels Vitrification of water by Rapid freezing of thin film suspensions

## Slide 25
tweezers
EM-grid
Liquid ethane
/propane
(-175 to -188 °C)
Native, aqueous sample is embedded in vitreous water and stored below -160°C
Micrograph(2D-projection)
Sample preparation for cryo-EM: vitrification of water (solvent)
Pipette 3ul sample
Plot excess liquid away
using filter paper
 Quench to the liquid ethane/propane

## Slide 26: Sample Preparation: Automatic system FEI Vitrobot™
Automate the cryo-fixation process at constant and user-definable physical and mechanical conditions e.g. temperature, relative humidity, blotting conditions (blot time, number of blots, blot pressure..)

## Slide 27
Notes: 27

## Slide 28
Notes: 28

## Slide 29: Cryo-transfer
Notes: 29

## Slide 30: Stantard cryo transfer holder... we have 5 holders: multiple specimen (3 samples) holder & cryo-rotation holder etc.
Resolution is limitted by sample holder due to thermal drift and vibrations - caused by external liquid nitrogen dewar.

## Slide 31: Dedicated cryo-TEM Liquid helium Jeol 3200FSC
Samples are surrounded by liquid helium/nitrogen cooling system inside the microscope - No external liquid nitrogen/helium dewar and samples are in small cartridges.
 Fast temperature stabilization after the sample loading
 Stable temperature – minimum thermal drift
 No vibration due to external devar
 True atomistic resolution (~2Å)
Eucentric tilt for tomography (±70 degrees)
Lattice resolution 0.204 nm at 18K and liquid nitrogen temperature

## Slide 32: Cryo-Stage
Liquid N2 or Liquid He
Liquid N2
cryo transfer holder for ”normal” TEM

## Slide 33
Liquid Helium
Liquid Nitrogen
Specimen
Specimen Exchange
Goniometer
Nitrogen Tank
Helium Tank

## Slide 34
Specimen

## Slide 35: Dedicated cryo-TEM  Stable specimen stage Very low sample Drift
~0.8 nm
~0.8 nm
Gold particels Magnification 1.2M
Same area after 6 min waiting
Specimen drift typically 1 - 2 nm/min (after 30 min and without pietzo), but when using the additional pietzo stage compensation – the drift can be as low as ~ 1 Å/min (as shown in the images above ~1.3 Å/min)

## Slide 36: Gold (200) lattice 0.204 nm
2.04Å

## Slide 37: Material Science Cryo-TEM Applications

## Slide 38: Example: vesicles in water.. Phosfolipid, polymeric etc.
Lipid bi-layers
Phosfolipid

## Slide 39: Example: Vesicles
Lipid bi-layers

## Slide 40: Biological hierarchical self-assemblied as templates for functionalities Native (Cellulose I) 3-10 nm nanofibers Exceptional properties Sustainable Large scale processes under developments (cf. silk)
Pääkkö, Ankerfors, Kosonen, Nykänen, Ahola, Österberg, Ruokolainen, Laine, Larsson, Lindström, Ikkala, Biomacromolecules 2007 8, 1934
Zimmermann et al, AdvEng Mat, 2004 6 754
Example: Cellulose Nanofibers

## Slide 41: Cryo TEM applications: Example: Cellulose Nanofibers in water solution
Dispersion of Individual 3-4 nm diameter fibers,
No need to remove the water, which may lead to fiber aggregation
One can directly see Nanofiber bundles in vitrified water solution.

## Slide 42: Liquid helium cryo TEM - Reducing the Beam damage

## Slide 43: Polylysine-lipid(DMPE-EG7) complex
SAXS
Notes: In this work we have used polylysine as the cationi polyelectrolyte. In the scattering curve of pure polylysine only a broad peak is observed. That corresponds to an average structure size of 15 Å. This is the diameter of pLys alfa-helix or one chain in a beta-sheet structure.
As the surfactant we have a triple tail lipid, which consists of two alkyl tails and one hydrophilic PEG tail. Here we have 14 methylene units. The pure lipid makes a lamellar structure with periodicity of 77 Å. In the lamellae the layers of hydrophobic alkyl tails and hydrophilic polyethylene oxide tails alternate.
The starting materials were combined in an aqueous solution in equimolar quatities and the resulting precipitate was dried. Here we can see the small-angle x-ray scattering curve fof the 1:1 complex. Unlike the pure lipid, the complex doesn’t show any easy lamellar structure. A more detailed analysis was needed to figure out the structure.

## Slide 44: Cryo-TEM at 86 K (Liquid nitrogen cooling )
SAXS
FFT from TEM
Low dose image taken at 86 K (- 187 oC)

## Slide 45: Cryo-TEM at 86 K (Liquid nitrogen cooling )
After one second
After few exposures ...

## Slide 46

Cryo-TEM at 18 K (Liquid helium cooling )

## Slide 47: Summary: Cryo-TEM Helium vs. Nitrogen
Helium
Nitrogen

## Slide 48: Cryo Vitrification: also other solvents possible
Toluene
Tetrahydrofuran (THF)
THF/methanol mixture
Alkohols (methanol, ethanol..)
..

## Slide 49: Example: Nanoparticles in solution (organic solvent)
Cryo electron microscopy by direct vitrification of solutions –
Benefit: There is no need to evaporate solvent – which may induce aggregation – but solvent is just vitrified in amorphous state.
How stable is the solution – is there any aggregation?
Cryo electron microscopy image after solution vitrification

## Slide 50: Example: rod-coil block copolymer aggregation in THF/methanol mixtures

## Slide 51
Fig. 7.6 Formation of HRTEM image. (a) Schematic imaging process
in TEM; (b) mathematical processing.
7.2 High-Resolution Transmission Electron Microscopy
7.2.1 Principles of HRTEM
Zhiping Luo, A Practical Guide to Transmission Electron Microscopy, Volume II Advanced Microscopy
CTF Lens Contrast Transfer Function

## Slide 52
CTF (Contrast Transfer Function) is the function which modulates the amplitudes and phases of the electron diffraction pattern formed in the back focal plane of the objective lens. It can be represented as:
Clearly, it will be a complicated curve which will depend on:
  Cs (the quality of objective lens defined by spherical aberration coefficient) , l (wave-length defined by accelerating voltage), Df (the defocus value), k (spatial frequency) (in figure u = k)
The contrast transfer function (CTF) is a mathematical description of the imaging process in the TEM, expressed in Fourier space. Ideally, TEM images would represent true projections of the electron density of the specimen. Instead, images are distorted by the microscope optics, representing primarily phase contrast rather than amplitude contrast, and contain high levels of noise
CTF Lens Contrast transfer function
Microscope resolution
In real space
Large features
Small features 
2 nm-1= 0.5 nm
4 nm-1= 0.25 nm

## Slide 53
CTF explorer demo

## Slide 54
2D FFT
Object
1D FFT would look a bit like this..
CTF
“final image” ~ FFT*CTF and then inverse FFT
Contrast will be very low in this focus setting
CTF Lens Contrast Transfer Function

## Slide 55
But when defocusing - CTF will move to left - so that the larger features has better contrast..
Above shown example case first two peaks would be perfectly transformed in normal contrast (i.e. electron dense areas are dark) , 3rd and peak information would disappear from the image and 4th and 5th peaks contrast would be inverse in the image
2D FFT
Object
CTF Lens Contrast Transfer Function
1D FFT would look a bit like this..

## Slide 56
Phase Plate technology in TEM’s: The Ultimate in Contrast – JEM-2200FS with Phase Plate!
Zernike phase plate in TEM: Thin carbon film with small hole in the center placed in the back focal plane (BFP) of objective lens – First models commercially available since ~ 2010-11
Ultramicroscopy, Volume 109, 2009, Pages 312–325
C

## Slide 57
The 300 kV cryo-TEM images of a sample from negatively stained horse spleen ferritin (a) ZPC-TEM image acquired using a Zernike phase plate. (b–d) Conventional TEM images at a defocus of 2550, 540 and 130 nm, respectively (underfocus). The insets show the diffractogram for each image. The scale bars in the insets correspond to 1 nm−1.
Kuniaki Nagayama, Journal of Electron Microscopy 60(Supplement 1): S43–S62 (2011) “Another 60 years in electron microscopy: development of phase-plate electron microscopy and biological applications”
defocus of 130 nm
defocus of 2500 nm
Good contrast 1.5 nm  0.26 nm
Good contrast 10 nm  ~2.5 nm
Good contrast > 0.4 nm

## Slide 58
Volta phase plate (VPP) which has similar but simplified design. It comprises a thin (~ 10 nm) continuous amorphous carbon film constantly heated to ~ 250oC. Unlike the Zernike phase plate the VPP does not have a central hole. The phase shift is created “on-the-fly” by the interaction of the central diffraction beam with the film. This characteristic makes the VPP very convenient to use because one does not have to precisely center it but just pre-irradiate an arbitrary position on the film.
Zernike thin film phase plate (ZPP) consists of a thin (~ 25 nm) material film with a small (~ 500 nm) hole in the center. It was the first phase plate to be used in cryo-EM applications and demonstrate the usefulness of phase plates.
ZPP plate had several practical disadvantages which prevented its wide adoption. The ZPP has a very short usable life of about a week after which it has to be replaced. It requires precise centering of the electron beam in the small central hole which makes automated data acquisition practically impossible. Last but not least, it produces fringe artefacts in the images due to the sudden onset of the phase contrast at the edge of the central hole.
Latest development is Volta phase plate (VPP)
Volta potential phase plate for in-focus phase contrast Transmission electron microscopy
Proc Natl Acad Sci U S A. 2014 Nov 4; 111(44): 15635–15640.