## Slide 1
SCHEDULE
Date Topic
1. Wed 01.03. Lec-1: Introduction
2. Mon 06.03. Lec-2: Crystal Chemistry & Tolerance parameter
3. Mon 06.03. EXERCISE 1
4. Wed 08.03. Lec-3: Crystal Chemistry & BVS
5. Fri 10.03. Lec-4: Symmetry & Point Groups
6. Mon 13.03. EXERCISE 2
7. Wed 15.03. Lec-5: Crystallography & Space Groups (Linda)
8. Fri 17.03. Lec-6: XRD & Reciprocal lattice (Linda)
9. Mon 20.03. EXERCISE 3 (Linda)
10. Fri 31.03. Lec-7: Rietveld (Linda)
11. Mon 03.04 EXERCISE 4: Rietveld (Linda)
12. Wed 12.04. Lec-8: ND & GI-XRD
13. Fri 14.04. Lec-9: XRR (Topias) at 12.15-14 (Ke4)
14. Mon 17.04. EXERCISE 5: XRR (Topias)
15. Wed 19.04. Lec-10: Synchrotron radiation & XAS & EXAFS
Fri 21.04. Mössbauer
Fri 21.04. EXERCISE 6
Fri 28.04. Seminars: EELS, XPS, FTIR, Raman
Wed 03.05. Seminars: ED, HRTEM, SEM, AFM
Wed 10.05. ADDITIONAL DISCUSSION/QUESTION POSSIBILITY
EXAM: Wednesday, May 17th

## Slide 2
LECTURE 9: XRR (X-ray reflection)
XRR complements GI-XRD in thin-film characterization
XRR: thickness, density, roughness & superstructures
KEYWORDS: diffraction, reflection & refraction, total reflection, critical angle, incident angle
Downloadall sizes Use this fileon the web Use this fileon a wiki Email a linkto this file Informationabout reusing

## Slide 3
DIFFRACTION (XRD & GI-XRD): Bulk phenomenon, takes place at the surface and also within the material, but only at certain angles. Bragg Equation: conditions and positions of interference peaks.
REFLECTION (XRR): Surface phenomenon. X-rays strike a surface and bounce back (reflect off it). Law of Reflection: reflection angle = incidence angle
REFRACTION: X-rays enter a different medium of different optical density and change direction or bend. Snell's Law: degree of refraction.

## Slide 4
GI-XRD
(1) Thin film + subs.
(2) Substrate
XRD
(3) Thin film + subs.
(4) Substrate
GI-XRD – last lecture
Small incident angle (α) lowers the penetration depth → information mostly from a thin surface layer
How θ & 2θ are changing in XRR measurement?

## Slide 5
Reflectivity of silicon
CRITICAL ANGLE
With very tiny incident angles X-rays only reflect from the sample surface (no refraction)
At the critical angle (αc), X-rays begin to penetrate (= refract) the material
Above αc the reflectivity rapidly drops
αc is material property → proportional to the square root of electronic density

## Slide 6
XRR (X-Ray Reflectivity)
Incident angle = Reflection angle: specular (mirror-like) reflection
No scattering from diffraction planes, only from surface/interfaces (where there is a change in refractive index)
Scattering depends on the properties of the two interface material layers
IDEALLY: scattering intensity depends on the layer thicknesses and electron densities of the two materials
IN PRACTISE: intensity depends also on surface (or interface) roughness

## Slide 7
XRR Measurement
XRR devise similar (same) to XRD device but the optical parts somewhat different
SCATTERING ANGLE fixed to the INCIDENT ANGLE
With the lowest incident angles, sample surface reflects all X-ray radiation intensity
When the incident angle increases above the critical angle X-rays start to penetrate to the sample (top layer)
When the incident angle is further increased the intensity starts to oscillate
Oscillation (so-called Kiessig fringes) is related to the mass densities and layer thicknesses
Oscillation provides us with important information of the sample
Typical measurement range: few degrees
(with higher angles the background noise level considerably increases)

## Slide 8
Information from XRR data
  You will see in the Exercise Session that the pattern can be simulated to get quantitative values, but even by eye observation, quite much information is obtained
  Critical angle (αc; total reflection limit ) → DENSITY
  Periodic oscillations or so-called Kiessig fringes provide us lot of information
  - Distance between two Kiessig fringes → THICKNESS
  - Decay rate of intensity → (SURFACE) ROUGHNESS
  - Hight of Kiessig fringes → (INTERFACE) ROUGHNESS
density
roughness
roughness and density contrast
thickness
Incident angle (deg.)
Intensity of reflection

## Slide 9
rel = 2.82 x 10-15 m
  Film Density
Higher the density, higher the critical angle
Amplitude of Kiessing fringes increases with increasing density difference (film vs. substrate)
   Stronger reflection/scattering

## Slide 10: EXAMPLE: ZnO thin film
Critical angle: 0.333 o

## Slide 11
300 nm thick film
is already too thick
for XRR
Film Thickness
Larger the distance between fringes,
smaller the film thickness
Measurable thickness range 1 – 200 nm (ideally)
For rough, crystalline samples maximum thickness ~100 nm or slightly above
Fringes “merge” together with higher film thickness
Roughness decreases the quality of the fringes
(La,Sr)CoO3 ALD thin films

## Slide 12
σ = 0 nm
σ = 1 nm
Thickness 60 nm
Smooth
Rough
Rough interfaces cause intensity loss
Surface roughness

## Slide 13
Interface roughness
Total intensity drop
Oscillation amplitude drop
 No clear film-substrate-interface for clean interference (oscillation)

## Slide 14
Superlattices  periodicity in XRR
ALD-grown superlattice: Eu2O3 and TiO2 layers grown on top of each other with different frequencies: in the XRR pattern you can see clear (more intense) superlattice peaks.
Superlayer thicknesses:
10N = 7.5 Å
20N = 15.9 Å
50N = 29.1 Å

## Slide 15
ELECTRON
PHONON
Organic
Oxide
J.-P. Niemelä, A.J. Karttunen & M. Karppinen,
Inorganic-organic superlattice thin films for thermoelectrics,
J. Mater. Chem. C 3, 10349 (2015).
EXAMPLE: THERMOELECTRIC MATERIALS
High electrical conductivity & Low thermal conductivity
→ Difficult combination to be achieved with conventional materials
ALD/MLD thin-film technology → nanoscale SUPERLATTICE (SL):
- thermoelectric oxide layers (ZnO) by ALD & organic (benzene) layers by MLD
- thermal conductivity decreases but electrical conductivity remains the same
XRR: we can see SL peaks as an indication of the regular ordered SL structure
XRR
Notes: 15

## Slide 16
N x [ZnO : benzene]
N
40
30
20
15
12
10
8
6
3
XRR:
We can see/count the number (N) of ”superlayer” units in the SL thin film; most clearly for N = 6 to 12; for N > 12 the oscillations start to overlap
NOTE: for ZnO no SL peaks are seen
T. Tynell, I. Terasaki, H. Yamauchi & M. Karppinen, J. Mater. Chem. A 1, 13619 (2013).

## Slide 17
F. Krahl, A. Giri, J.A. Tomko, T. Tynell, P.E. Hopkins & M. Karppinen, Thermal conductivity reduction at inorganic-organic interfaces: from regular superlattices to irregular gradient layer sequences, Adv. Mater. Interfaces 5, 1701692 (2018).
Superlattice
Gradient film
Using the ALD/MLD technique it is possible to perfectly control where within the ZnO film the organic (benzene) layers are placed → We can grow both regular superlattice films and irregular “gradient” ZnO-organic films. In this EXAMPLE we compare two films:
In both: total film thickness: ~105 nm
Number of organic layers: 5
Average ZnO layer thickness: ~17 nm
Superlattice: all ZnO layers ~17 nm
Gradient film: ZnO layers 9 ~ 28 nm
ONLY for the former the SL peaks are seen in XRR data
XRR data
Gradient film
Superlattice

## Slide 18
A. Philip et al., J. Mater. Chem. C, 2022,10, 294-300
EXAMPLE: ε-Fe2O3:azobenzene superlattices